MSACL-CAC_Batch Cover Sheet template_for plates and ALH Extraction

Sample Table/.xxx file name_________________________________	LCMS_1	LCMS_2

Extraction CLS____ Data Processing CLS____ Plate A1 placement/Smpl-Rck barcode  _____

See reverse and reagent logs for reagent, blank, I.S. and calibrator lot #s.

1. Scanned plate barcode_________________  2.1st plate use/wells today_100824_85-96

3. Record Heat Block Temp if used ______	TEST1: acceptable range 63 to 67°C

Note sample set-up exceptions here:

Data Review:									Notes:

Observed Rts of all of Highest Calibrator I.S. peaks verified against S.O.P.

- Peak Shapes acceptable			______    ______________________________________

- RRts and Rts acceptable			______    ______________________________________

- Baselines acceptable				______    ______________________________________

- S/N of Low Calibrator acceptable		______    ______________________________________

- I.S. Peak areas acceptable			______    ______________________________________

- Cal Curve R2 and % dev. acceptable    	______    ______________________________________

- MRM ratios OK when ≥ LLOQ/cutoff  	______    ______________________________________

- Check for carryover			 	______    ______________________________________

- All blanks used for Neg QC acceptable 	______    ______________________________________

Manual integration performed		Y          N

Record Corrective Actions:

- # of re-injected samples			_______

- # of repeated samples                                _______

- Prevent reporting of results to be RPTd _______

- Test specific rework recorded                  _______

- QC acceptable	 				Y          N

- QC entered into QC software		Y          N

- Batch verified in LIS (except for repeats)

manually         interface

- If interface used – Instrument Menu reconciled

with Repeat Log

- Batch printed:

to file                 hard copy

20. Review batch for (test specific exceptions)

| Assay | Material | Lot# | Expiration Date | Check or Substitute |
| --- | --- | --- | --- | --- |
| AAA | Working Internal Standard | 09072024 | 12-07-24 |  |
|  | Calibrator | 10182024 | 12-18-24 |  |
|  | Extraction Reagent | 101824 | 12-19-24 |  |
|  | Wash Reagent (50 mmol/L pH 9) | 102024 | 12-13-24 |  |
|  | Elution Reagent (check pH <2) | 110924 | 12-14-24 |  |
|  | Extraction Plate | 2016050301 |  |  |
|  | Injection Plate | 12010640 |  |  |
|  | Injection Plate Cap | 139330-2 |  |  |
|  | ALH 350 uL nested tips | D154822I |  |  |